## LinkML_slots
classes | slot_group | slots | name | range | minimum | maximum | regex | multivalued | inlined_as_list | enums | description | is_identifier | recommended | required | slot_uri(ontology) | prefix_reference | comments
Flow cytometry experiment | identifiers | investigation_identifier | Investigation_ID | string | [6 empty] | The identifier for this investigation collected using the "investigation" template. This is used to match the flow cytometry experiment to an investigation. |  |  | TRUE | obo:OBI_0000066 | http://purl.obolibrary.org/obo/ | SRP_PRJ1
Flow cytometry experiment | identifiers | study_identifier | Study_ID | string | [6 empty] | The identifier for this study collected using the template appropriate for your study design (e.g., MIATE for animal experiments). This is used to match the flow cytometry experiment to a study. |  |  | TRUE |  |  | study_001
Flow cytometry experiment | identifiers | sample_identifier | Sample_ID | string | [6 empty] | The identifier for this study collected using the template appropriate for your study design (e.g., MIATE for animal experiments). This is used to match the sample source to the data. |  |  | TRUE |  |  | HSA_001
Flow cytometry experiment | identifiers | flow_experiment_identifier | Flow_experiment_ID | string | [6 empty] | This is a unique ID assigned by the user to represent each individual record in this flow cytometry experiment. | TRUE |  | TRUE |  |  | flow_001-HSA_001
Flow cytometry experiment | experiment | purpose |  | string | [6 empty] | A brief description of the goal of the experiment. This should include the rationale and hypothesis. |  | TRUE |  | obo:NCIT_C25634 | http://purl.obolibrary.org/obo/ | Test the hypothesis that CD5+ innate-like B Cells (ILBs) are sensitive to AHR-mediated immunotoxicity.
Flow cytometry experiment | experiment | keywords |  | string |  |  |  | TRUE | TRUE |  | The keywords should include terms from an appropriate vocabulary (e.g. MeSH) to describe the experiment. Each keyword should be seperated by a ";" with not spaces. |  | TRUE |  | obo:NCIT_C43513 | http://purl.obolibrary.org/obo/ | Monocytes;2,3,7,8-Tetrachlorodibenzo-p-dioxin
Flow cytometry experiment | experiment | experimentVariables |  | enum |  |  |  | TRUE |  | Condition;Dose;Timepoint;Individual;Experimental variable sample type | Variables are attribute(s) that differ between samples within an experiment due to pre-existing differences in sample states or due to experimental manipulation of the samples. |  |  | TRUE | obo:STATO_0000258 | http://purl.obolibrary.org/obo/ | Dose;Timepoint;Individual
Flow cytometry experiment | experiment | primaryContact |  | string |  |  |  | TRUE |  |  | The individual who is the primary contact for this experiment (last name,first name). |  |  | TRUE | obo:OBI_0000202 | http://purl.obolibrary.org/obo/ | Blevins,Lance
Flow cytometry experiment | experiment | date |  | date | [6 empty] | The date or time period during which the investigation was performed (i.e., from collecting and treating samples to performing data analysis) shall be stated. |  | TRUE |  | obo:STATO_0000093 | http://purl.obolibrary.org/obo/
Flow cytometry experiment | experiment | conclusions |  | string | [6 empty] | A brief summary of the interpretation of the results or outcome of the experiment shall be provided if available. |  | TRUE |  | obo:IAO_0000144 | http://purl.obolibrary.org/obo/ | Levels of AHR-mediated suppression of the IgM response from individual donors strongly correlated with the percentage of the B cell pool that was CD5+, suggesting that CD5+ B cells are more sensitive to AHR-mediated impairment.
Flow cytometry experiment | experiment | qualityControlMeasures |  | enum |  |  |  | TRUE |  | STR profiling of source cells;mycoplasma testing;isotype control for antibody;no antibody negative control;positive control cell type;Live/Dead staining;unstained controls | A description of the quality control measures used such as replicates, calibrations, control assays, etc. shall be provided. If another experiment was performed for the purposes of establishing quality control standards, that experiment may be referenced. |  | TRUE | TRUE | obo:MSIO_0000085 | http://purl.obolibrary.org/obo/ | Live/Dead staining;unstained controls
Flow cytometry experiment | experiment | flow_cytometry_protocol |  | string | [6 empty] | A brief description of the flow cytometry protocol |  |  |  | obo:NCIT_C42651 | http://purl.obolibrary.org/obo/ | Cells were processed using a BD FACSCanto II using FACS Diva software (BD Bioscience) and analyzed using FloJo (Version 10, Treestar Software, Ashland OR).
Flow cytometry experiment | experiment | flow_cytometry_protocol_DOI |  | string | [6 empty] | A DOI to a version controlled protocol, preferably available on a repository such as protocols.io |  |  |  | obo:OBI_0002110 | http://purl.obolibrary.org/obo/
Flow cytometry experiment | Instrument | flowCellType |  | enum | [5 empty] | jet-in-air;quartz cuvette | The flow cell type (e.g., stream-in-air, cuvette) including the material shall be provided. |  |  |  | obo:OBI_0400043 | http://purl.obolibrary.org/obo/
Flow cytometry experiment | Instrument | other_flow_cell |  | string | [6 empty] | Other relevant flow cell details should be provided and may include flow cell manufacturer name (e.g., MIFlowCyt Optics, Inc.) and model number (e.g., PRO-CFC-75), flow cell outside shape (e.g., rectangular with an attached sphere, spherical, square) and dimensions, flow channel shape (e.g., rectangular, circular, triangular) and dimensions, sheath flow velocity (e.g., 10 l/s), sheath flow pump type (e.g., gas pressure, vacuum, gear), sample flow velocity (e.g., 7.5 l/s) and sample flow pump (e.g., gas pressure, syringe). |  |  |  | obo:NCIT_C17649 | http://purl.obolibrary.org/obo/
Flow cytometry experiment | Instrument | manufacturer |  | enum | [5 empty] | BD Biosciences | The instrument manufacturer shall be specified by stating its name (e.g., MIFlowCyt Cytometry Systems, Inc.). A URL pointing to manufacturer web pages may also be provided. |  |  |  | obo:OBI_0000835 | http://purl.obolibrary.org/obo/ | BD Bioscience
Flow cytometry experiment | Instrument | model |  | enum | [5 empty] | BD FACSAria | The instrument model shall be specified (e.g., CytMeter Elite-2000). It may include the instrument serial number. |  |  |  | obo:MS_1000031 | http://purl.obolibrary.org/obo/ | BD FACSAria
Flow cytometry experiment | Instrument | configurationAndSettings |  | enum | [5 empty] | The instrument has not been altered | The system and its configurations have significant impact on experimental results and as such the following shall be provided |  |  |  | obo:MS_1000740 | http://purl.obolibrary.org/obo/ | The instrument has not been altered
Flow cytometry experiment | Instrument | other_instrument |  | string | [6 empty] | Additional relevant instrument details should be provided and may include machine-specific information such as information about automated agitation, temperature control, controlled volume dispensing, sampling from microtiter plates, auto-boost, auto-flush, etc. Other custom settings shall be provided if relevant and may include setting name, description, and value. |  |  |  | obo:NCIT_C17649 | http://purl.obolibrary.org/obo/
Flow cytometry experiment | light_source_instrument | type |  | enum | [5 empty] | mercury-vapor lamp; xenon lamp;diode laser; dye laser; gas laser;solid state laser;light-emitting diode;tungsten halogen lamp | The type of the light source shall be provided (e.g., laser, diode laser, xenon lamp). |  |  |  | bao:BAO_0150029 | http://www.bioassayontology.org/bao/
Flow cytometry experiment | light_source_instrument | excitatoryWavelength |  | string | [6 empty] | The excitatory wavelength shall be provided (e.g., 488nm, 633 nm). |  |  |  | bao:BAO_0000566 | http://www.bioassayontology.org/bao/
Flow cytometry experiment | light_source_instrument | powerAtExcitatoryWavelength |  | string | [6 empty] | The power of the light source shall be provided (e.g., 180 mW, 50 mW). If possible, this should be measured at the intersection of the light source beam with particles; see also 3.3.3. |  |  |  | bao:BAO_0000566 | http://www.bioassayontology.org/bao/
Flow cytometry experiment | light_source_instrument | polarization |  | string | [6 empty] | Polarization shall be specified (e.g., linear, >100:1, in the direction of the sample flow). If possible, this should be specified at the intersection of the light source beam and particles; see also 3.3.3. |  |  |  | obo:OMIT_0009913 | http://purl.obolibrary.org/obo/
Flow cytometry experiment | light_source_instrument | beam |  | string | [6 empty] | Light source beam dimensions and geometry shall be specified at the flow cell or at the intersection of the beam with particles (e.g., elliptical 22 x 66 Î¼m). |  |  |  | obo:NCIT_C70569 | http://purl.obolibrary.org/obo/
Flow cytometry experiment | light_source_instrument | other_light_source_instrument |  | string | [6 empty] | Additional relevant light source details should be provided and may include light source manufacturer name (e.g., MIFlowCyt Lasers, Inc.), light source model name (e.g., 302C Krypton-Ion Laser), light source part number (e.g., I-326), noise (e.g., <1%), pointing and power stability (e.g., <1% over 1 hour), etc. |  |  |  | obo:NCIT_C17649 | http://purl.obolibrary.org/obo/
Flow cytometry experiment | optical_detector | name |  | string | [6 empty] | The optical detector name (e.g., FSC, SSC, FL1, FL2, FL3, FL4, etc.) shall be specified.
Flow cytometry experiment | optical_detector | type_optical_detector |  | enum | [5 empty] | photodiode;photomultiplier tube | The type of the optical detector shall be provided (e.g., photodiode, photomultiplier tube). |  |  |  | obo:FBbi_00000269 | http://purl.obolibrary.org/obo/
Flow cytometry experiment | optical_detector | voltage |  | float | 0 | [5 empty] | The optical detector voltage shall be specified (e.g., 300V, 700V).
Flow cytometry experiment | optical_detector | amplificationType |  | string | [5 empty] | linear;log | The optical detector amplification type shall be specified (e.g., linear, log). Additional details including the actual amplifier profile may be provided.
Flow cytometry experiment | optical_detector | amplificationFactor |  | float | [6 empty] | For log amplification, the number of decades shall be provided (e.g., 4 decades). For a linear amplification, the linear gain shall be provided (e.g., linear gain of 2.0).
Flow cytometry experiment | optical_detector | other_optical_detector |  | string | [6 empty] | Additional relevant optical detector details should be provided and may include the optical detector manufacturer name and model number (e.g., MIFlowCyt Photonics Inc., PMT #R9220), minimal, maximal, and peak wavelength sensitivity (e.g., 185 nm, 900 nm, and 650 nm), the internal gain (e.g., 107), or the rise time (e.g., 2.2ns).
Flow cytometry experiment | optical_filter | type_optical_filter |  | enum | [5 empty] | band pass filter;long pass filter; dichroic long pass filters;short pass filter;dichroic short pass filter; polarizer excitation filter; grating, prism | The optical filter type shall be specified.
Flow cytometry experiment | optical_filter | transmittedWavelengths |  | string | [6 empty] | Light wavelengths transmitted by the filter shall be specified (e.g., 488/25 nm, >670 nm, <620 nm).
Flow cytometry experiment | optical_filter | installationDate |  | date | [6 empty] | As optical filters of all types are subject to degradation, filter performance needs to be monitored at intervals to verify continued performance at an acceptable level [19]. The optical filter installation date shall be specified (e.g., January 15, 2007).
Flow cytometry experiment | optical_filter | other_optical_filter |  | string | [6 empty] | Other relevant optical filter details should be provided.
Flow cytometry experiment | flourescence_reagent | staining_protocol |  | string | [6 empty] | A brief description of the flow cytometry protocol |  |  |  | obo:NCIT_C42651 | http://purl.obolibrary.org/obo/ | ~0.25x10^6 cells were stained with antibody.
Flow cytometry experiment | flourescence_reagent | staining_protocol_DOI |  | string | [6 empty] | A DOI to a version controlled protocol, preferably available on a repository such as protocols.io |  |  |  | obo:OBI_0002110 | http://purl.obolibrary.org/obo/
Flow cytometry experiment | flourescence_reagent | characteristicsBeingMeasured |  | string | [6 empty] | The relative amount of molecules, properties, or processes being evaluated (e.g., CD25, apoptosis, membrane permeability, cell viability, oxidative burst). The characteristic(s) being measured shall be provided whenever there is ambiguity about the analyte being measured, such as when sample processing affects detection of the analyte (e.g., propidium iodide example in Table 1). The optical detector (e.g., FL1) or parameter (e.g., FL1-H, see 3.3.6) used primarily for this measurement shall be indicated. |  | TRUE |  | obo:BFO_0000005 | http://purl.obolibrary.org/obo/ | CD5
Flow cytometry experiment | flourescence_reagent | analyte |  | string | [6 empty] | What plays the role of the analyte? Which substance or chemical constituent is the subject of interest of the analytical procedure, what target is specifically bound by the analyte detector (e.g., CD25)? |  |  | TRUE | obo:OBI_0000275 | http://purl.obolibrary.org/obo/ | CD5
Flow cytometry experiment | flourescence_reagent | analyteDetector |  | string | [6 empty] | What plays the role of analyte detector? Which component of the fluorescence reagent specifically binds to the analyte to make it detectable (e.g., anti-CD25 antibody)? |  |  | TRUE | obo:CHMO_0002244 | http://purl.obolibrary.org/obo/ | mouse anti-human CD5
Flow cytometry experiment | flourescence_reagent | analyteReporter |  | string | [6 empty] | (Fluorochrome) What plays the role of analyte reporter? Which component of the fluorescence reagent reports the presents of the analyte to the flow cytometer? What substance (label) is used to generate the measured signal (e.g., FITC)? |  |  | TRUE | obo:CHEBI_35209 | http://purl.obolibrary.org/obo/ | PE
Flow cytometry experiment | flourescence_reagent | cloneNameOrNumber |  | string | [6 empty] | If the probe is a monoclonal antibody, the clone name or number shall be provided. |  |  |  | obo:SO_0000151 | http://purl.obolibrary.org/obo/ | UCHT2
Flow cytometry experiment | flourescence_reagent | reagentManufacturerName |  | string | [6 empty] | The reagent manufacturer shall be specified (e.g., MIFlowCyt Reagent Inc.). |  |  |  | obo:IAO_0000590 | http://purl.obolibrary.org/obo/ | BioLegend
Flow cytometry experiment | flourescence_reagent | reagentCatalogueNumber |  | string | [6 empty] | The reagent catalogue number shall be specified. |  |  |  | obo:NCIT_C112406 | http://purl.obolibrary.org/obo/ | 300607
Flow cytometry experiment | flourescence_reagent | stainingStatus |  | enum | [5 empty] | stained;unstained | [6 empty] | stained
Flow cytometry experiment | flourescence_reagent | other_fluorescence_reagent |  | string | [6 empty] | When secondary antibodies are used, the primary analyte (the analyte bound by the primary antibody) and the reporter component of the secondary antibody (the reporter that generates the measured signal) shall be clearly specified as described in 2.4.2, 2.4.3, and 2.4.4. If relevant, the primary and secondary antibodies shall be described individually (e.g., in the case that the use of a secondary antibody significantly impacts the interpretation of results because of nonspecific fluorescence, etc.). The description should also contain other relevant details which may include lot number, concentration, label incorporation method, separation technique, antibody staining procedures, and wash steps. |  |  |  | obo:NCIT_C17649 | http://purl.obolibrary.org/obo/
Flow cytometry experiment | data_analysis | analysis_protocol |  | string | [6 empty] | A brief description of the flow cytometry protocol |  |  |  | obo:NCIT_C42651 | http://purl.obolibrary.org/obo/ | Positive gates were drawn using unstained and day 0 controls.
Flow cytometry experiment | data_analysis | analysis_protocol_DOI |  | string | [6 empty] | A DOI to a version controlled protocol, preferably available on a repository such as protocols.io |  |  |  | obo:OBI_0002110 | http://purl.obolibrary.org/obo/
Flow cytometry experiment | data_analysis | listModeDataFile |  | string | [6 empty] | The list-mode data files (e.g., FCS files [18]) shall be provided directly or details on how they may be requested shall be stated. |  |  |  | obo:OBI:0000021 | http://purl.obolibrary.org/obo/ | sample1.FCS
Flow cytometry experiment | data_analysis | CompensationDescription |  | string
Flow cytometry experiment | data_analysis | TransformationPurpose |  | string | [6 empty] | The purpose of each performed transformation shall be specified (e.g., data visualization, background correction, statistical analysis, quantitative flow cytometry, etc.).
Flow cytometry experiment | data_analysis | TransformationDescription |  | string | [6 empty] | Either the exact mathematical formulas/algorithms of each data transformation shall be supplied using an open and freely available specification, or a description of each transformation shall be provided.
Flow cytometry experiment | data_analysis | GateDescription |  | string | [6 empty] | The subpopulation identified by the gate shall be briefly described (e.g., “IL-4 producing helper T cells of the CD3+CD4+ phenotype”). | [5 empty] | ILBs of the CD5+ phenotype
Flow cytometry experiment | data_analysis | Gate Statistics |  | string | [6 empty] | Percentage of events within the gate shall be provided specifically stating the denominator. The denominator shall be either the total population of events (e.g., percentage of lymphocytes based on the total number of events) or another gate (e.g., percentage of CD4+ lymphocytes based on all lymphocytes). When the denominator is another gate, this shall also be exactly specified in recursive fashion so that the gating strategy can be followed up to the original data set. The denominator may not necessarily be a containing (superset) population for some statistics (e.g., the donor / host blood cell ratio).
Flow cytometry experiment | data_analysis | GateBoundaries |  | string | [6 empty] | Either the exact mathematical descriptions of each gate boundary shall be provided using an open and freely available specification, or this information shall be provided in the form of images.
Flow cytometry experiment | data_analysis | WorkspaceFile |  | string | [12 empty] | Analysis.wsp
Flow cytometry experiment | provenance | template version |  | provenance
Note on A1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: STRING Classes are used to separate a standard into different DataHarmonizer tables. For example: Investigation, Study, Assay could be represented by three different templates.
Note on B1: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: STRING These are used to organize the template into sections. For example: Animal details, housing conditions, substance exposure, ...

## LinkML_description
id | name | description | version | license
https://example.com/MIFlowCyt | flow_cytometry_assay | Template based on the expectactions from the FlowRepository (flowrepository.org) and the Minimum Information about Flow Cytometry Experiments (MIFlowCyt) developed by the International Society for Advancement of Cytometry (ISAC). More details can be found at https://doi.org/10.25504/FAIRsharing.kcnjj2. The LinkML Schema and template was developed by the Michigan State University Superfund Research Center Data Management and Analysis Core. | 1.0.0 | CC0